Ge 116 Winter 2009
Lab project Week 1: Scanning Electron Microscope, part 1 (instrument use and imaging)

Report due Tuesday January 20. Electronic submission is acceptable, so is paper.

This exercise is a combination of learning to use the instrument and actually analyzing your sample. Many steps are things to do rather than questions to answer. Products that should be turned in are shown in boldface. Make yourself a directory on the SEM computer to store your results.

A. Required reading for instrument check-out

All new SEM users are required to read the following three documents:

http://www.gps.caltech.edu/facilities/analytical/sem.html
http://www.gps.caltech.edu/facilities/analytical/sem_user.html
http://www.gps.caltech.edu/facilities/analytical/sem/LEOSEM_CIT_v2.pdf
Please read these before your lab session this week.

A. Sample preparation- probably not necessary as samples were prepared last year, but we will check each section at the beginning of lab.
You will receive your sample with most of the preparation work already done, i.e. as a 1” round polished section. However, for high-quality SEM work some additional sample preparation is needed.

1. Final vibro-polish: learn to use the vibrating polisher in the Analytical Facilities lab and give your sample about 1 hour of final polishing in colloidal silica suspension. This will give a finish good enough the EBSD work next week. Clean the vibro-polisher when you’re done. 

2.
Carbon coat: learn to use to the carbon coater in the Analytical Facilities lab and give your sample a thin carbon coat (suitable for SEM use; we will recoat it with a thick coat for Electron Probe analysis in two weeks). Clean the carbon coater when you’re done.

B. Basic instrument operation

1. Start-up, vacuum system, sample loading: Start the SEM software, sign-in. Get your sample ready to mount on the SEM stage. Put some identifying marks (C-tape) on the surface of your sample, this will aid in consistent orientation from instrument to instrument.  Vent the chamber. Insert your sample. Pump to vacuum (learn about vacuum system diagnostics, high vacuum and variable pressure modes). For now, use high vacuum mode. Learn about sample chamber layout, chamber camera viewing, basic instrument safety.

2. Electron column, beam conditioning: Learn how to get an electron beam, choose accelerating potential and beam current, align the electron optics, focus the beam at given working distance, and stigmate the beam. For starters, use 10 kV accelerating potential, a 30 m aperture (do not select high-current mode), and an 8 mm working distance.

3. Stage and scanning controls: In the course of focusing, you will learn the steps necessary to find and focus on regions of interest on your sample: stage navigation, scanning (scan rates and raster sizes), and zooming.

C. Electron imaging

1. Secondary electron image: Secondary electrons are most sensitive to sample topography. Our SEM has several secondary electron detectors; learn the strengths and weaknesses of each. We will verify that our sample is flat and well-polished and that the beam condition steps above are correctly done by taking SEI images. Capture a survey SEI image of your sample at lowest possible magnification. Choose an imperfection on the surface (a pit or crack) and zoom in on it. Capture the highest magnification SEI image of this feature at which you can get good image quality (~ 1000x).

2. Backscattered electron image: Operating conditions: 15 keV acc. voltage, 120 um aperture, 8 mm W.D. Backscattered electrons are most sensitive to mean atomic number of the sample material. With optimized contrast it is easy to distinguish variations between adjacent elements in the periodic table at moderate concentrations. There are trade-offs between imaging parameters (accelerating potential, aperture, working distance) that we will explore somewhat. Our instrument has two backscattered electron detectors; learn about the strengths and weaknesses of each. Go back to low magnification, optimize brightness and contrast to distinguish among the minerals in your sample and using Stage Control capture a series of BEI images that can be tiled to make a navigation map of the general mineral texture of your whole section. Be sure not to change the brightness or contrast from one image to the next in your mosaic; you want medium gray in one image to be medium gray in the next… Some students find it helpful to record the size of their grid (it can help in making the mosaic).  Pick 3 areas of interest and capture moderate magnification BEI images of these areas.  Try to note the location of these areas.  We will return to these areas over the course of the quarter; you can pick areas that seem to have most of your major minerals, have an abundance of trace minerals, an area with neat BEI patterns or zoning, etc.  Now do the following steps at least three times on different major mineral targets: zoom in on a single grain of some mineral, optimize brightness and contrast to reveal any zoning that might be present within the grain. Design imaging parameters for maximum Z contrast and capture a moderate-magnification BEI of the grain. Finally, zoom in on a single accessory mineral or inclusion, design imaging parameters for maximum spatial resolution and try to obtain the highest magnification BEI image at which you can get good image quality.

3. When you are done, retrieve your sample, store it in a labeled container in the dessicator in the SEM lab, and return the machine to standby properly. Transfer your image data to yourself (by FTP, e-mail, memory stick, whatever).

D. Interpretation (i.e. work on these questions after you are done using the machine)

1. Mineralogy: Based on backscatter contrast, how many different minerals can you identify in your sample? Rank the minerals by typical backscatter brightness (noting that some may be zoned and cover a range of brightnesses) and characterize each distinct mineral by typical shape and size (estimate the mean size for each phase). You may at this stage apply petrological insight or hand-sample observation to attempt to identify some of the minerals, but we will get more diagnostic information later so this would just be preliminary.

2. Texture: Based on your BEI navigation map, describe the texture of the rock. Is there a shape-preferred orientation apparent on the plane of your section? Defined by what textural features? Are there any other notable textures, such as variations in size, overgrowth or replacement textures, symplectites, etc? Given that you may not yet have identification of the minerals that make up the texture, summarize the textural features of your sample (mineral size, anisotropy, relations between minerals, etc.).

3. Areas of Interest: On a copy of your BEI map, locate and mark the 3 areas of interest.  Bring a copy of your BEI map, marked with areas of interest, as well as BEI images of the areas themselves to the next lab session.  

