## Slide 1: Ge 116: introduction
This is a class in analytical methods, but what does that mean?
  Operating the instruments, obviously, but more importantly…
  Understanding the physical basis of each kind of measurement and
  Understanding the fundamental statistical nature of measurement, the nature of uncertainty, the difference between accuracy and precision, and systematic vs. random errors.

## Slide 2: Ge 116 Module 1: Scanning Electron Microscopy
Part 1: Electron optics, beam-specimen interactions & imaging

## Slide 3: What is an SEM?
It is not a TEM (transmission electron microscope)
  TEM is the electron-optics equivalent of a transmission optical microscope, taking advantage of the tiny deBroglie wavelength of electrons to achieve ultra-high spatial resolution, with tiny field of view and need for ultra-thin sample foils
It is not a STEM (scanning tunneling electron microscope) or AFM (atomic force microscope)
  These instruments put a probe extremely close to a sample and move the sample the build up a raster image, depending on near-field interactions to determine sample properties at atomic resolution

## Slide 4: What is an SEM?
Rather, it is a Scanning Electron Microscope
  A well-focused beam of electrons is scanned in a raster (or boustrophodon) pattern over an area on the surface of a stationary sample
  Signals generated by beam-specimen interactions are measured as a function of time during the scan and mapped to spatial position on the sample
  Images are built up by plotting signal intensity vs. position — like a conventional cathode-ray tube television. Noise is reduced by scanning slowly or by integrating over many scans
  The scan can be stopped to focus temporarily on a point of interest

## Slide 6: Parts of an SEM
• Vacuum chamber and pumps
Electron source
Electron column:
  High Voltage
  Condenser, Objective, and Stigmating optics
  Scan coils
Stage
Detectors
Analog or digital display

## Slide 7: Electron sources: Filament

## Slide 8: Electron sources: Filament
“Work function”
Fermi level
Ground state

## Slide 9: Electron sources: Field Emission

## Slide 10: Electron sources: Field Emission

## Slide 11: Electron sources: Comparison
Source | Brightness (A/cm2sr) | Lifetime | Source Size | Energy Spread | Beam Stability
W filament | 105 | <100 h | >30 m | 2.5 eV | 1%
LaB6 | 106 | <1000 h | >5 m | 1.0 eV | 1%
Schottky Field Emission | 108 | >1000 h | < 30 nm | 0.75 eV | 2% in 2005, 1% now
Cold Field Emission | 1010 | >3000 h | <1 nm | 0.35 eV | 2%
JEOL claims total currents up to 2 mA and 1%/hr stability for their new Schottky FE electron gun

## Slide 12: Electron optics
Condenser Lenses
Apertures
  Together these determine beam current
Stigmators
  Determine beam shape
Scan Coils
  Determine magnification and scan rate
Objective Lens
  Determines working distance/focus

## Slide 13: Electron Beam Limits
Minimum beam diameter
~1
Spherical Aberration Coefficient, ~2 cm
Electon DeBroglie wavelength, ~9 pm
Probe current
Source Brightness
Understanding the trade-off between current and resolution
Note, however, that beam diameter is not the same thing as imaging or analysis resolution…for many reasons.

## Slide 14: Electron Beam Limits

## Slide 15: Beam-Specimen Interactions
Elastic
  Back-scattered electrons
Inelastic
  Secondary electrons
  Bremsstrahlung X-rays
  Characteristic X-rays
  Auger photoelectrons
Other
  Cathodoluminescence
  Specimen current

## Slide 16: Beam-Specimen Interaction Volume

## Slide 17: Beam-Specimen Interaction Volume
Increased backscatter yield for inclined specimen surface
Horizontal spread of backscatter emission for normal and inclined specimens
Locations of secondary electron emission

## Slide 18: Backscattered Electrons

## Slide 19: Backscattered Electrons
(b is specimen current image)

## Slide 20: Secondary Electrons

## Slide 21: Secondary Electrons

## Slide 22: Design of image acquisition
What spatial scale of feature(s) are you trying to see?
  Relative to excitation radius achievable with beam diameter, source, and instrument optics at a given current and voltage
  Always focus and stigmate as well as possible
What contrast mechanism will make feature(s) visible?
What scan rate will give sufficient signal to noise ratio?